## Slide 1: LM Analysis – compare old and new PMTs –60 hrs Dataset Runs 15920 to 15990 22/4 - 1:09 pm to 25/4 – 2:20 am

## Slide 2: Old and New PMTs
Calorimeter | Old PMT | New PMT
11 | 25 | 5
12 | 26 | 0
13 | 27 | 6
18 | 28 | 8
20 | 29 | 4
Looking at calo 20 in detail – overall behaviour of all channels in back up. Since the amplitude is free from jitter of the pedestals, I used the amplitude instead of area for this presentation (though area plots are the same). Details of amplitude of first pulse here, as it will have the effect of temperature of the PMT alone.

## Slide 3: Long term stability / fluctuation
“Forward” component of signal which goes to calos
“Return” component which comes back from calos
Correction to be applied only if the “return” signal changes. To check this use old and new PMTs.
Red – signal of the New PMT
Green – signal of Old PMT
R = A2:A1 of new PMT and R’ = A’2:A’1 of old PMT. In the ratio of these ratios could indicate the stability of the return signal i.e. R: R’ (new:old)
Courtesy: A. Fioretti, C. Gabbanini, and G. Venanzoni

## Slide 4
LM (after discussion with INO Colleagues)
WFD g-2 (scheda “Corradi”)
Local DAQ (scheda “Napoli”) & eventualmente ai WFD g-2
Fibra PMMA
Fibra quarzo
Foward
Backward
A2=L*F*B, F= forward eff; B=Bacw eff
IF F Gain Sipm should be corrected; IF B no
R=A2/A1  most of the effect due to F  apply to Gain corr
DR=R_PM1/R_PM2 sensitive to B1/B2 (no to Gain corr)
A2

## Slide 5: Comparison of old and new PMTs for Calo 20
Normalized area first pulse
Normalized area second pulse
Old PMTs about 1.5 times greater drop than new ones

## Slide 6: Comparison of old and new PMTs for Calo 20
New PMT – Calo 20, channel 4
Old PMT – Calo 20, channel 29
Bin size = 50 mins.

## Slide 7: New PMT – channel 4 for Calo 20
Normalized area first and second pulse overlaid
Normalized A2:A1 for old and new PMT

## Slide 8
Ratio of ratios for Calo 20
R = A2:A1 of new PMT and R’ = A’2:A’1 of old PMT. The ratio of these ratios R: R’ (new:old) is not very stable => The “return” component is not stable.

## Slide 9: Temperature Variation – Channel 4
Note: The temperature plots changed due to some additional calibration corrections applied. This resulted in re analysis and and different correlation of temperature with gain.
Laser hut temperature
Optical table temperature
Overlaid temperature and amp 1 on the same plot

## Slide 10: Channel 4 New PMT – Temperature vs. Amp1
Optical Table Temperature
Behaviour of each time slice. A simple graph appears to show similar trend for each part (anticorrelated with temperature).

## Slide 11: Amplitude Vs. Temp
The slope appears to change at 29.8 C.
Amp 1
A2:A1
Amp 2

## Slide 12: Back Up Slides

## Slide 13: Other data sets
11th April - 30 hours
22nd April - 30 hours

## Slide 14: Temperature variation of Ratio of ratios

## Slide 15: A2:A1 /oC fpr all channels

## Slide 16
Some Studies/Observations –Scatter plot and projection of a bin – good Gaussian

## Slide 17
Correlation between Ped / amplitude ratio